## Slide 1
Electrons capture in the air
The subject:
Medical Ionization Chambers are air filled ICs
In the air, electrons are captured by O2
 forming O2- (this is just the beginning of the story…)
When increasing dose rate, we increase recombination rate
X+ + Y-  X + Y
Recombination depends on the time ions spend mixing
it is useful to know how many electrons (fast)
are captured to form O2- (slow).
How to perform this measurement a direct way?

## Slide 2
Electrons capture in the air
Before answering, let’s play with Ramo Theorem…
5mm
+HV
Virtual ground
+
-
Rf
50W
Rg
Simple wide-band line receiver
We will see later
Incident a particle
Suppose we do something like that… What would we get?
Nitrogen atm. P
Voltage gain = 50

## Slide 3
Electrons capture in the air
Number of charges created by incident particle (5.5MeV a)
a in Nitrogen
in
5.5 MeV
4.27cm
3.67cm
out
5.0 MeV
DE= 0.5MeV
What : electrons/ions pairs
Where : uniformly along the trajectory

## Slide 4
Electrons capture in the air
Electric field in the detector :
+HV
0V
Th
Solve :
Z
z=0
To obtain :
This time, that was easy, isnt’it?
Surprises will come later…

## Slide 5
Electrons capture in the air
Secondaries transport :
+HV
0V
Th
Z
z=0
electrons
ions
Are moving
At velocity
For instance
Th=6mm
HV=1000V
E=1666V/cm
Mobility (#constant)
Same E  same velocity everywhere

## Slide 6
Electrons capture in the air
Secondaries transport : not a cumbersome problem…
 make it by hand!
For electrons:
time
charges
N
Full collection
time
Dt # 360ns

## Slide 7
Electrons capture in the air
Ramo virtual field :
0V
1V
Th
Remove all space charges
Put 1V on measuring electrode
Put 0V on others electrodes
Z
z=0
Intensity :
Uniform E field

## Slide 8
Electrons capture in the air
Current generator :
0V
1V
Th
z=0
For electrons:
instrumentation examples
8
time
Current (A)
Full collection
time
Dt # 360ns
imax # 6.2 nA

## Slide 9
Electrons capture in the air
And total charge :
For electrons:
time
Current (A)
Full collection
time
Funny, isn’t it?

## Slide 10
Electrons capture in the air
And now, show time!
We don’t want to distort signals… which are very very small!
+HV
0V
5.5 MeV a source
Plastic scintillator
PM tube
FASTER D.A&P.S.
start
Signal
It’s Ad time: you need a numerical state of the art, general purpose
Digital Acquisition System? Have a look @ faster.in2p3.fr

## Slide 11
Electrons capture in the air
And now, show time!
Acquire each signal triggered by PM & average them
We want to measure 50x50x6.2nA # 15 µV
Noise # 600µVRMS
With 1Mhits :
Signal is always 15 µV
Noise # 600µVRMS / 1M½ = 0.6µVRMS

## Slide 12
Electrons capture in the air
And now, show time!
Dt # 360ns
imax # 6.2 nA
Detector HV = 1kV
Energy loss =13900 charges
Line receiver preamp.
Rl = 50W
Voltage gain = 50
Vmax # 15 µV

## Slide 13
Electrons capture in the air
What happens in the air?
Free electrons are captured at constant rate :
The shape of current becomes :
For a given HV :

## Slide 14
Electrons capture in the air
The show goes on :

## Slide 15
Electrons capture in the air
Ramo, it works!
air
nitrogen
Look at the beginning :
Currently admitted values
 | air | nitrogen
w (eV) | 35.1 | 36.4
S/r (MeV.cm2/g) | 7.12 102 | 7.21 102
r (g/cm3) | 1.205 10-3 | 1.165 10-3
I prop to | 2.44 | 2.31
Ratio  1.06

## Slide 16
Electrons capture in the air
That’s all folks!
This work was performed by Guillaume Boissonnat (Thx G. !)
Ramo-Shockley theorem is a wonderful tool for
signal prediction
It describes exactly what happens in your detector
If you don’t observe what you predicted
 your model (detector, physics) is false
 your detector doesn’t work properly