## Slide 1
PM signals quantification.
…Toward a good methodology
A PM tube is a vacuum device
Photocathode
Converts incident
Photons into
Photoelectrons (PE)
dynodes
Amplifies the
Number of PE
Anode
 You get the signal there
3 photons  1PE
1PE  1M electrons

## Slide 2
PM signals quantification.
…Toward a good methodology
We bought a PM (or more plausible, found it in a cupboard…),
is it the reason to believe the constructor?
How to check the gain?
How to check the optical coating & coupling we took hours to do?
or…
or…
or…
Is it still working?

## Slide 3
PM signals quantification.
…Toward a good methodology
PM : XP2020
HT = 1650V
Single photo-electron response
PM
Input impedance
Rload = 50W
The signal we see thru
the scope is completely
related to Ramo
current generator!
 We can use it to measure some detectors properties

## Slide 4
Tr # 2ns
Tf # 3ns
A # 2mV
PM signals quantification.
…Toward a good methodology
Mean measured amplitude
For 1 P.E. :
And remember :
What the scope shows :
imax # 2mV/50W = 40µA
Q = imax *(Tr + Tf)/2 = 100fC
Gain = Q / e = 0.625 106
From detector to measurement
From measurement to detector
1 P.E ↔ 5mV.ns

## Slide 5
PM signals quantification.
…Toward a good methodology
NaI (Tl)
This is not (electronics) noise
Just the aleatory time emission of
photons by the scintillator
40mV
10mV
450ns
500ns
Surf = 40 x 450/2 + 10 x 500/2 = 11500 mV.ns
 2300 P.E.
g 662keV

## Slide 6
PM signals quantification.
…Toward a good methodology
PM Quantum efficiency :
CsI (Tl)
BGO
NaI (Tl) & most organic scintillators
LaBr
#30%
Qeff
bialkali
P.E. = Qeff x photons

## Slide 7
PM signals quantification.
…Toward a good methodology
So, I measured 2300 P.E. for 662keV
I should have
w is about 40eV for NaI(Tl)
Qeff # 0.3 for my XP2020
Collection efficiency # 46%
+

## Slide 8
PM signals quantification.
…Toward a good methodology
Stupid, I am!!!
2Rpk # 4,2cm
2Rpm # 5,25cm
Useful Area / PM Area = 0,64 !!!
NaI(Tl)
XP2020
photocathode

## Slide 9
PM signals quantification.
…Toward a good methodology
Chart: 
NaI (Tl)
BC400
LaBr
662keV
#470keV
Note: spectra have
been shifted of a decade
for clarity
# 46keVrms
# 9keVrms
aucun
Compton Plateau
Tot. Abs. Peak

## Slide 10
PM signals quantification.
…Toward a good methodology
It is easy, and often useful, to check detectors basic properties
LaBr is cool, what a pity it’s so expensive…
Btw, somebody broke mine, does someone have 5000€ to give me? Nop?
In an instrumentation chain, some things are not calculable
(for instance, here, the collection efficiency)
But, with the help of simple measurements, it’s always
possible and instructive to infer these (human) parameters.
It explains how our understanding of the chain is far (or not)
from it’s real performances
That’s all folks!